On the characterization of Plutonium distribution in MIMAS MOX by image analyZis

Ghislain Oudinet, Isabelle Munoz-Viallard*, Marie-Jeanne Gotta

Commissariat à l’Énergie Atomique (CEA), Département d’Étude des Combustibles - CEA/ Cadarache, F-13108 Saint-Paul-Lez-Durance, France

J.M. Becker

CPE (Chimie Physique Électronique) Lyon Bat. 308, 43 bd du 11 novembre 1918, F-69616 Villeurbanne CEDEX, FRANCE

G. Chiarelli,  R. Castelli

COGEMA, Business unit recycling - 30204 Bagnols sur Cèze CEDEX -FRANCE.

*  Corresponding author. Fax : 33-4 42 25 47 17 ; e-mail : isabelle.munoz-viallard@cea.fr

Abstract

A better understanding of MOX fuel in-pile behaviour requires a very detailed characterization of the Pu distribution in the pellet before and after irradiation. Electron Microprobe Analysis (EPMA) can be used to determine chemical elements distributions with a 1 µm spatial resolution. This paper describes the development of X-ray microanalysis techniques to produce semi-quantitative “maps” of plutonium concentrations in order to characterize, in a short time, and with reasonable accuracy, large areas of fuel microstructure (1 mm2). A new segmentation technique is then proposed, based on statistical compatibility, so as to finely describe the MIMAS MOX fuel microstructure. To demonstrate the reliability of this new method, three materials are finely characterized. In each case, the results demonstrate the good and reliable accuracy of this new characterization methodology. The analysis method used is currently able to identify three so-called phases (the Pu-rich agglomerates, a reticulated phase and the uranium-rich agglomerates) and to quantify the plutonium distribution and the plutonium content of these three phases. The impact of the fabrication process on the microstructure is seen in the surface distribution variations of the plutonium and in the local plutonium content variations.

# 1. Introduction

MOX fuel has been used in French PWR since 1987. In October 2005 about 2400 MOX fuel assemblies (1100 tM of MOX) have been loaded in 20 different EDF 900 MWe reactors. Maximum MOX assembly burn-up is currently 42 GWd/tM, but several experimental rods have reached a burnup close to 60 GWd/tM.

As UO2 fuel is allowed to operate up to 52 GWd/tM, this burnup stands as the first objective for MOX fuel in order to reach what we call “MOX parity” (same management and burn up for UO2 and MOX), planned to start in 2007 [0]. Increasing the performance of MOX fuels relies on knowledge and understanding of the in –pile behaviour of the current MOX fuel. Irradiated fuel surveillance programs conducted since the beginning of plutonium recycling in French PWRs have shown that fission gas release is higher in MOX fuel rods than in UO2 ones, at the same burnup [1], [2], [3].

This higher gas release in MOX fuels has clearly two reasons and possibly a third one:

- for core physics reasons, the power density, hence the linear heat rate decreases at a slower rate in MOX fuels than in UO2 fuels,

- the thermal conductivity of MOX fuel is slightly lower than the UO2 one : the centerline temperature is therefore higher even at the same power [4]

- the microstucture of MOX pellets fabricated by the MIMAS process (MIcromized MASter blend) where a primary blend of UO2 and PuO2 is diluted with UO2, exhibits a non homogeneous distribution of plutonium. Is this special microstructure partly responsible for higher gas release ? Some studies have shown differences in gas releases for MOX fuels elaborated with different manufacturing process and/or different UO2 [5], [6].

By a detailed investigation of the evolution of the fuel microstructure and of fission gas behaviour, it is now possible to draw a more comprehensive picture of the mechanisms related to gas release and of the influence of the initial microstructure on these mechanisms. Fission gas release models and codes that have been developed by CEA [7] are able to integrate MOX microstructure effect, in order to improve the modelling of in pile behaviour of MOX fuel. These models use as input data the characterizations of as-fabricated MOX fuel and they use for validation the characterization of irradiated MOX.

This paper explains the methods developed to characterize initial MOX fuel microstructures and especially the plutonium distribution within the fuel, using microprobe images. This information is very important because local fission density is directly related to the way fissile plutonium atoms are initially distributed in the material.

First, the automated method of semi-quantification of X-ray mappings associated to the plutonium element is described. Then, the method of automatic segmentation into several compounds (typical of MOX fuel), based on the counting statistic associated to this image acquisition technique, is shown. All these methods are then applied to two different industrial fuels.

# 2. Presentation of the analyzed material

The MIMAS (MIcronized MASter blend) process is today the most widely used industrial process for MOX fuel. It consists of two distinct steps. A mixed powder made out of plutonium dioxide, uranium dioxide and scrap, with a resulting plutonium content lower than 30%, is finely ground. The obtained powder, called “master blend”, is then sifted and diluted in uranium dioxide in order to reach the desired plutonium content. This secondary mixture is then pressed into pellets, which are then sintered and centerless ground to the specified diameter.

Uranium oxide and plutonium oxide can form solid solutions (U,Pu)O2, but the apparent interdiffusion coefficients of U and Pu are very slow under reducing conditions (  m2/s at 2000K and kJ/mole) [8] [9]. Sintering manufactory conditions are unfavourable for homogenisation and for the formation of solid solutions of UO2-PuO2 mixtures.

The manufactory process results in a material consisting in Pu-rich agglomerates with a plutonium content close to that of the master blend, and of U-rich agglomerates without plutonium. These agglomerates are separated by a coating whose plutonium content is in between. We will thereafter use the word “phase”, not in the sense usually intended in chemistry, but to distinguish the fractions of the material with the different plutonium concentrations.

The studied materials A and B are industrial fuels worked out by COGEMA - MELOX

- A is a typical standard MOX MIMAS ; it is a high plutonium content pellet (7.2% Pu/(U+Pu)) randomly sampled out of a MELOX batch,

- B is an experimental MOX MIMAS fuel with a high plutonium content (7.1% Pu/(U+Pu)) processed so that the plutonium agglomerates are smaller than the standard ones [10] [11].	
On the microscopic scale, the X-ray mappings acquired with an electronic microprobe allow the fine analysis of the coating phase in which UO2 and PuO2 form a continuous solid solution (FIG. 1, where darker pixels correspond to low plutonium concentration).

A : standard MIMAS MOX fuel       B : optimised MIMAS MOX fuel

180 µm

FIG. 1. X ray images of Pu in non-irradiated standard MOX (materials A and B).

A universal characterization method is required for this step to objectively compare the different batches and to quantify the evolution of their microstructure during irradiation.

First, P. Garcia and all [12] have performed microstructural analysis, using microprobe examinations on MOX MIMAS fuels manufactured by Belgonucléaire. X ray mappings were analysed using the common method of histogram thresholding. The threshold was defined as the grey level at which the second derivative of the grey scale histogram reaches a local maximum. Our aim was to look for more robust method, not too sensitive to the operator effect and to the image noise.

Image analysis is an efficient tool, which has enabled the development of automated methods of semi-quantification and of segmentation of microprobe mappings [13] [14]. These methods are described below.

# 3. Experimental methods

The experimental method was performed in three stages:

- the acquisition of X-ray mapping of plutonium,

- the quantification of the image;

- the segmentation of the image.

The originality of this study is based upon the technique of semi-quantification of the image and on the method of segmentation, carried out starting from a semi-quantified image.

### 3.1 -  Micropobe images acquisition

The number of  X mappings to be acquired depends on the homogeneity on the manufacture batch. Typically, in the framework of this study, a longitudinal section and a cross section were produced. On each metallographic cut, three 1 mm2 in area fields (resolution : 1024x1024 pixels) were selected randomly by means of the displacement mode of the sample-holder stage. Measurements were carried out on the peaks Pu, U and O without subtracting the background noise of the spectrum (continuous background). The select counting time is close to 20 ms per pixel, which accounts for approximately 6 hours of acquisition by field.

### 3.2 – Method of semi-quantification

3.2.1 –Acquisition of a quantitative line

Each X mapping  is connected to a quantitative profile taken randomly on the image. Such profiles are carried out by steps of 1 µm in size. Intensities measured on the M line of Pu, on the standard and samples to be analyzed, are corrected considering the background noise and removing the interference between the plutonium M line and the uranium M line. Accurate plutonium concentration is then measured starting from the relative intensity between the analyzed sample and the standard by means of ZAF-type correction factors. Counting duration over each point lasts 15 s for peaks and 10 s for the continuous background. The standards used are pure UO2 for uranium and oxygen, and of pure PuO2 for plutonium. The acceleration voltage of  electrons and the intensity of the probe current selected for the quantitative analyses are respectively equal to 15 kV and 80 nA.

#### 3.2.2- Calibration line

The association of X mappings of Pu with quantitative profiles makes it possible to establish a calibration line giving the Pu content versus counts per nA and per sec (grey level) acquired for each pixel of the image. With this objective, the first stage consists in extracting the grey level profile corresponding exactly to the quantitative profile analyzed on the image. However, precise positioning on X mappings of the points of the profile to be quantified (materialized by the pixels of the image) depends on the accuracy of the displacement of the sample holder stage, as well as on slight fluctuations acquisition conditions. Thus, an algorithm of correlation is used to precisely  locate this profile on the image (FIG. 2). The coefficients of correlation obtained are greater than 0.95. The equation translating grey levels of the image to plutonium concentration can then be established by means of a linear least squares regression such as those described in details by Ancey et al. [15]. The image in level of grey is transformed into a semi-quantified image.

FIG. 2. Agreement between a quantitative profile and grey levels of pixels extracted
from an electronic microprobe image (laboratory sample).

### 3.3 – Segmentation method

Traditional methods of thresholding of grey level histograms (maximization of the variance between classes [16, 17], maximization of entropy [18, 19], methods using various order derivatives of the histogram [20], etc.) turned out not to be very reliable and too sensitive to the image background noise. The specific method of thresholding proposed within the framework of this study integrates simultaneously grey levels and the vicinity of each pixel. In addition, it takes into account the statistics of counting associated with X-ray emission making it possible to determine a confidence interval for each measured mean value. Counting associated with each measurement point follows a statistical model of distribution of Poisson type. Consequently, starting from the semi-quantified image, zones statistically compatible with the initial content of the master blend (representing plutonium rich agglomerates) can be defined as well as those compatible with null plutonium content (representing uranium agglomerates). Remaining zones are subjected to a more complex analysis, their plutonium content being intermediate between that of plutonium rich zones and that of uranium rich ones.

A geodetic distance to the nearest plutonium agglomerate is associated to each point of the image. This distance is the length of the shortest path connecting this point to a plutonium agglomerate, circumventing uranium agglomerates. X mappings are then transformed into a geodetic chart of distances, where the level of gray of each point not belonging to a plutonium agglomerate is proportional to the shortest distance separating it from a plutonium agglomerate. This chart enables us to study the average plutonium concentration as a function of the distance to the nearest plutonium-bearing cluster. The layout of the graph (cf FIG. 3) representing the average plutonium concentration of a point of the image according to its distance to the nearest plutonium agglomerate (among two studied fuels) is the cornerstone of the method of segmentation suggested hereafter.

This graph highlights a major reduction in plutonium concentration within the first micrometers surrounding clusters and a stabilization at long distances. Fluctuations at the end of the graph come from a low statistical representativeness (few points of the image are located at such a long distance from plutonium  agglomerates). The decrease as well as the stabilization of the plutonium content show respectively two definitely linear profiles. The graph can thus be modelled by two segments : one directed almost vertically, the other directed almost horizontally.

FIG. 3. Mean Pu concentration versus the nearest Pu-rich agglomerate distance

The maximum concentration, obtained at a zero distance from plutonium agglomerates, is characteristic of the minimal content of the agglomerates core. We regard the right part of the graph (approaching the horizontal one) as characteristic of the average content of the coating phase. The level of thresholding in concentration is thus to be located between the low limit value associated with the master blend, and the mean level of concentration associated with the coating phase (intersection between the two right-hand side modelled segments). In the absence of any additional information, the level of thresholding is taken halfway between these two levels of concentration (see figure 3). Uranium agglomerates cannot undergo the same treatment insofar as their grey levels include a considerable amount of background noise. Thus they are simply assimilated to their core on statistical compatibility grounds. FIG. 4 illustrates the segmentation of the image presented in FIG. 1, for sample A. The three phases are clearly identified. The morphology of uranium agglomerates (black zones) and of plutonium agglomerates (white zones) is respected. The coating phase (gray zone) appears to be connected, which is coherent on a physical point of view.

FIG. 4. Segmentation result (material A).

The interest of this method is primarily based on consideration of the vicinity of each pixel of the image, making it possible on the one hand to dissociate the pixels belonging to the master blend of a local over-concentration of fine plutonium particles, and on the other hand, to adapt to any type of morphology of agglomerates.

### 3.4 – Qualification of the data processing method

Our quantification method was qualified by comparing the results obtained during precise quantitative measurements using a microprobe on a 2500 points grid, to those obtained in the same field using the microprobe qualitative image and our translating equation from grey levels to concentrations. The results obtained on part of a U-rich agglomerate and part of a Pu-rich agglomerate (Table 1) shows plutonium quantification results of part of a microprobe image of a MOX fuel.

Table 1. Qualification of the quantification method (laboratory sample)

| Pu % | Pu-rich agglomerate | U-rich agglomerate |
| --- | --- | --- |
| Microprobe quantification | 36.09(0.01 | 0.19(0.01 |
| Calculated | 36.2(0.2 | 0.00(0.12 |

NB.: This qualification is based on samples other than the A and B. Especially, the initial plutonium concentration of the master blend in this sample is 40%.

# 4. Experimental results on fresh fuel

## 4.1. Detection of phases

The result of the phase segmentation differentiates three phases within the MOX MIMAS fuel and determines the surface fraction of each of them. The surface fraction of the phases is shown in Table 2 for the analyzed materials. The uncertainties associated to the calculated values correspond to the scatter among the six analyzed fields per batch (confidence interval at two times the standard deviation). These are comparable from one batch to another. The volume percentage of the master blend introduced during the fabrication phase is close to 26%. The proportionality factor between the amount of master blend introduced during fabrication and the surface fraction of agglomerates measured by image analysis is noted Fp. This factor is greater than 1 since a noticeable part of the plutonium is transferred in the coating phase during the subsequent fabrication steps. It is close to 2 in the case of batch A and greater than 3 in the case of batch B (Table 2). The difference noticed on the Fp factor shows the influence of the grinding-sifting process. The plutonium and uranium phases are less extensive in batch B, where the coating phase is predominant. This is illustrated in the microprobe image shown in FIG. 2, sample B.

Table 2. Phases distribution and proportionality factor Fp.

|  | Material A | Material B |
| --- | --- | --- |
| Pu-rich agglomerate (Area %) | 14 ( 2 | 7 ( 1 |
| “Coating phase” (Area %) | 51 ( 1 | 61 ( 1 |
| U-rich agglomerate (Area %) | 36 ( 1 | 31 ( 1 |
| Fp | 1.9 | 3.5 |
| Total Pu content (%) | 7.2 | 7.1 |

## 4.2. Size spectrum of Pu-rich agglomerates

Within the framework of the R&D program of MOX fuels, the aim was to reduce the size of the plutonium enriched agglomerates [11]. This parameter may influence the retention and/or release of fission gases especially during off-normal events, although this effect has not yet been fully quantified [5, 6]. The comparison between batches A and B shows that with a similar final Pu content, the size spectrum of the Pu-rich agglomerates has changed towards smaller sizes (FIG.5). The population of small agglomerates (< 20 µm) occupies a surface fraction that is twice as large ; about 29% of the total agglomerate surface, to be compared to 12% in the case of batch A. One can notice on the cumulated graph that in batch B, more plutonium is contained into small size agglomerates, which was the expected goal.

FIG. 5. Pu-rich  agglomerate size spectra, and cumulated plutonium percentage relative to the total plutonium content.

## 4.3. Size spectrum of U-rich agglomerates

The size spectrum of uranium agglomerates is illustrated in FIG. 6. The granulometric distribution of the initial UO2 rough powder batches is today centred around 30 µm. The sintering shrinkage (about 15% of the volume), the small U/Pu global interdiffusion, as well as a possible erosion of the agglomerates during the dilution phase in the fabrication process, tend to move this spectrum towards slightly smaller sizes. FIG. 6 shows distributions centred around 20 µm. The average calculated sizes for the U-rich agglomerates are respectively 27 µm and 20 µm.

FIG.6. U-rich agglomerate size spectra.

## 4.4. Global plutonium content

The quantification of the microprobe images allows access to the average plutonium content of the analyzed fields and to the average plutonium content of each phase. Table 3 illustrates the differences observed between the calculated values and the experimental values measured by chemical dissolution of (U, Pu)O2 pellets. Good agreement is obtained between the calculated values using the quantification method and those expected.

Table 3. Global Pu content.

| Pu/(U+Pu) content (wt %) | Material A | Material B |
| --- | --- | --- |
| Calculated | 7.2 ( 0.4 | 7.4 ( 0.3 |
| Experimental | 7.2 | 7.1 |

## 4.5. Analysis of different phases: quantification of their distribution and plutonium concentration

The surface distribution is analyzed as a function of the three phases : plutonium rich agglomerates, coating phase and uranium agglomerates. An element of comparison was defined by the Fp factor. This factor takes different values and indicates that for the same fabrication process, the final content in the Pu-rich agglomerates is not directly proportional to the volume fraction of master blend introduced.

For the batch B, the percentage of Pu-rich agglomerates is reduced and the fine particles are spread in the “coating phase” which then becomes the largest one. The quantity of agglomerates coming from the master blend and their morphological characteristics (cohesion degree, granulometry, density…) have a direct impact on the plutonium distribution in the final sintered pellet.

The mean plutonium content in each of the phases has been quantified (Table 4). The two batches have master blend agglomerates with an equivalent plutonium content (about 24% Pu/(U+Pu)O2).

Table 4. Phases Pu content.

| Pu content (wt %) | Material A | Material B |
| --- | --- | --- |
| Pu-rich agglomerate | 24.0 ( 0.2 | 23.9 ( 0.2 |
| “Coating phase” | 5.7 ( 0.1 | 7.1 ( 0.2 |
| U-rich agglomerate | 0.5 ( 0.1 | 0.5 ( 0.1 |

It is worth noting that for batch B, much more than half of the plutonium is contained in the coating phase (67 % total Pu) (Table 5). In a standard MOX fuel the plutonium is equally divided between the two phases : agglomerates and coating. The microprobe coupled with the image analysis treatment allowed a better understanding of the microstructure of MOX fuels, on the microscopic scale.

Table 5. Pu distribution inside the phases of the three materials tested.

| Pu/Pu total % | Material A | Material B |
| --- | --- | --- |
| Pu-rich agglomerate | 51 ( 4 | 30 ( 2 |
| “Coating phase” | 46 ( 3 | 67 ( 2 |
| U-rich agglomerate | 3 ( 1 | 3 ( 1 |

# 5. Use of the segmentation method on a two cycles irradiated fuel

## 5.1. Segmentation

The fuel irradiation in reactor has two main consequences:

— the plutonium is burned by fission reactions that locally reduce its concentration and adds fission products, much lighter,

— the neutronic fertilization of U238 increases the plutonium content in the uranium zones.

Within a fuel pellet, the progressive averaging of the plutonium concentration during irradiation makes the identification of plutonium areas more difficult. It is however of major importance to be able to analyze in the same way the fresh fuel and the irradiated fuel, so as to characterize its evolution.

The segmentation method presented above being based on the notion of homogeneous domain, we decided to test it on fuel irradiated for two cycles.

## 5.2. Experimental results

The analyzed material was a two cycles irradiated MOX MIMAS fuel. The average content of the pellet (4.86 % Pu/(U+Pu)) was determined by microprobe analysis. A quantitative profile of 126 points (1µm per point) as well as a qualitative mapping (area of 1024x1024 pixels with 1µm2 per pixel and grey levels ranging from 1 to 199) was made. The semi-quantification of the image of the material led to a positioning of the profile on the map with a correlation coefficient of 0.93 (FIG. 7). The average calculated value of the field is then 4.62 % Pu/(U+Pu), a value consistent with the interfiled variations by a few tenths of percent.

Our segmentation method is based on the knowledge of the plutonium content in the “core” of the agglomerates, as well as nil content in plutonium in the “core” of uranium agglomerates. These assumptions are not verified in the case of irradiated fuel. This is why the plutonium content in the “cores” of the different agglomerates (Pu-rich, and U-rich) was evaluated using recordings found on the semi-quantified map.

Following the semi-quantification and segmentation phases, the agglomerates and the coating phase are visually contoured correctly and in a similar way to what is obtained on fresh fuel (FIG. 8).

FIG. 7. Agreement between the Pu concentration and the grey levels on an image of a 2 cycles irradiated fuel.

FIG. 8. Application of the segmentation method to an image of a two cycles irradiated fuel.

The segmentation method would require more validation tests on irradiated fuel. However, its operating principle as well as first results indicate that the characterization of two cycles irradiated fuels is of equivalent quality to that obtained on fresh fuel. The lack of matter (pores, cracks) which appear during irradiation are easily identifiable, which is of considerable interest in finding a reliable characterization of irradiated fuels.

# 6. Conclusion and outlook

The industrial objective to increase the burn-up of MOX fuels in France first to 52 GWd/tM and then beyond this value requires further analysis and understanding of the behaviour of MOX fuel under irradiation. It is thus necessary to set up and qualify tools of in-depth characterization of the microstructures. The development and validation on fresh and irradiated fuel of a method based on an automatic threshold of the images acquired with a Castaing electronic microprobe allow the quantification of the different phases (UO2, Pu rich agglomerates and coating phase) and the improved comparison between the different types of MOX microstructures, during their whole lifetime. Quantified elements are given for the distribution and plutonium content of the three detected phases. They show a noticeable impact of the fabrication process. The goal to obtain a more homogeneous Pu distribution with Pu-rich agglomerates of reduced size has been reached.

Bibliography

[0] J.L. Provost, M. Debes MOX and UOX PWR Fuel Performances – EDF operating experience. Water Reactor Fuel Performance Meeting, Kyoto, October 2005.

[1] L. Brunel, P. Blanpain, G. Chaigne, M. Trotabas "MOX fuel performance in French reactors : recent results and improvement programmes", International Symposium on MOX fuel cycle Tecnologies for medium and Long Term Deployment, Vienna Austria, May, 1999

[2] P. Blanpain, X. Thibault, M. Trotabas "MOX Experience in French Power Plants" ANS International Topical Meeting on Light Water Reactor Fuel Performance, West Palm Beach, april 1994

[3] P. Blanpain, X. Thibault, PJ. Pagès "Recent results from the in reactor MOX fuel performance in France and improvement program" ANS International Topical Meeting on Light Water Reactor Fuel Performance, Portland, Oregon, march 1997

[4] L. Caillot, G. Delette, JP. Piron, C. Lemaignan, A. Chotard, JP. Berton "Analytical studies of MOX fuel behavior in reactors" ANS International Topical Meeting on Light Water Reactor Fuel Performance, Portland, Oregon, march 1997

[5] Y. Guerin, J. Noirot, D. Lespiaux, C. Struzik, P. Garcia, P. Blanpain, G. Chaigne, “Microstructure evolution and in reactor behavior of MOX Fuel”, Proceeding of the International Topical Meeting on light reactor performances. Park City, USA, April 2000

[6] S.B. Fischer, R. J. White, P.M.A. Cook, S. Bremier, R.C. Corcoran, R. Stratton, C.T. Walker, P.K. Ivison and I.D. Palmer “Microstructure of irradiated SBR MOX fuel and its relationship to fission gas release”, J. Nucl. Mat., 306, (2002), 153-172.

[7] C. Struzik, M. Moyne, JP. Piron "High burnup modelling of UO2 and MOX fuel with METEOR/TU version 1.5." ANS International Topical Meeting on Light Water Reactor Fuel Performance, Portland, Oregon, march 1997

[8] S. Pillon, S. Mendez, D. Warin, R. Lorenzelli "Cation Diffusion and Crystal Growth at the UO2-PuO2 Interface",  Proceedings of the AIEA Technical Committee Meeting on Recycling of Plutonium and Uranium in Water Reactor Fuel, 3-7 July 1995, Newby Bridge, UK. IAEA-TECDOC-941.

[9] S. Mendez "Etude de l’interdiffusion U-Pu appliquée au combustible MOX", these de l’Université d’Aix Marseille, juillet 1995

[10] P. Blanpain, C. Callens, W. Goll, G. Chiarelli, JL. Guillet “MOX Fuel Performance and developpement” TOP FUEL 2001, Stockholm, May 2001.

[11] Ch. Delafoy, P. Blanpain, S. Lansiart, Ph Dehaudt, G. Chiarelli  “Advanced PWR Fuels for high burnup extension and PCI constraint elimination” AIEA Technical meeting on improved fuel pellet materials and design, Brussels, Belgium, 20-24 october 2003

[12] P. Garcia, A. Bouloré, Y. Guérin, M. Trotabas, P. Goeriot “In pile densification of MOX fuels in relation to their initial microstructure”, Proceeding of the International Topical Meeting on light reactor performances. Park City, USA, April 2000

[13] Gh. Oudinet, I. Munoz-Viallard, MJ. Gotta, JM. Becker, G. Chiarelli, R. Castelli  “On the characterization of plutonium distribution in MIMAS MOX by image analysis” AIEA Technical meeting on improved fuel pellet materials and design, Brussels, Belgium, 20-24 october 2003.

[14] G. Oudinet, " Analyse d’images et modélisation 2D/3D de la microstructure du combustible MOX ”, Thèse de l’Université Jean Monnet, Saint-Étienne, March 2003.

[15] M. Ancey, F. Bastenaire, R. Tixier, J. Phys. D 10, pp 817-830, 1977.

[16] M. Coster, J.L. Chermant "Précis d'analyse d'images", Presse du CNRS, March 1989.

[17] F. Charollais, "Analyse d'images : un outil pour caractériser et modéliser la microstructure du combustible MOX", Thèse de l'université de Caen, October 1997.

[18] J. N. Kapur, P.K. Sahoo, A.K.C. Wong, "A new method of gray level picture thresholding using the entropy of the histogram", Computer vision, graphics and image processing, vol. 29, pp 273-285, 1985.

[19] R. Zeboudj, "Filtrage, seuillage automatique, contraste et contours : du prétraitement à l'analyse d'images", Thèse de l'université de Saint-Etienne, April 1988.

[20] A. Bouloré " Etude et modélisation de la densification en pile des oxydes nucléaires UO2 et MOX ", thèse à l’Institut National Polytechnique de Grenoble et de l’Ecole Nationale Supérieure de Mines de Saint Etienne, March 2001.

# Comments

Jusqu'à l'année dernière, Fra prétendait qu'il n'était pas besoin de prendre en compte la microstructure dans les modèles de FGR

il vaut mieux ne pas donner de valeurs ici puisqu'on en donne dans la section 4.1 qui sont d'ailleurs différentes de celles-ci.

Il faudrait peut-être tracer les segments sur la figure

pourquoi la figure est-elle construite à l'inverse de ce qu'on a l'habitude d'utiliser ?

ce n'est pas montré sur la figure

ce n'est pas vraiment le même

pour éviter la confusion avec uranium enrichi

voir commentaire 5